## Slide 1
PRESENTED BY:
CLETUS NGWERUME & FREDERICK DAMPTEY
WRIST FRACTURE DETECTION USING CONVOLUTIONAL NEURAL NETWORKS
ARTIFICIAL INTELLIGENCE IN HEEALTHCARE
GROUP 9
SAT 5114

## Slide 2
Wrist fractures are among the most common orthopedic injuries, especially in elderly and trauma patients.
  Wrist joint comprises multiple small bones and articulations, making diagnosis complex.
  Plain X-rays are cost-effective and widely used but require expert interpretation.
  Issue: Subtle fractures are often missed, especially in understaffed or rural healthcare settings.
AI can assist clinicians by providing rapid, consistent interpretations — reducing diagnosis time and improving outcomes.
INTRODUCTION
A Computer-Aided Diagnostic system for accurate and early wrist fracture detection using deep learning models

## Slide 3
Problem: Human expert readers like radiologists and orthopedic surgeons are not always reachable to interpret x-rays. Even when available, subtle fractures may easily be missed, resulting in misdiagnosis. This is compounded by variability in interpretation.
Project Goal: Develop a deep learning-based Computer-Aided Diagnostic (CAD) system using CNNs.
Objectives:
  Acquire, load, preprocess and classify wrist X-ray images (fracture vs. normal).
  Train ResNet50 and EfficientNetB0 using transfer learning.
  Evaluate performance using accuracy, Precision ,F1-score, AUC.
  Use Grad-CAM to visualize and localize fractures on X-rays.
PROBLEM STATEMENT & OBJECTIVES

## Slide 4
CNNs in radiography have shown strong performance in prior studies.
Suen et al. (2024): AI system achieved 92% sensitivity, 89% specificity in wrist fracture detection.
Husarek et al. (2024): Meta-analysis found CNNs achieved pooled sensitivity/specificity ~90%.
Yee et al. (2019): Inception-ResNet v2 with Faster R-CNN gave an AUC of 0.895.
LITERATURE REVIEW

## Slide 5
Source: Kaggle (Better Mura)
Size: 59171 images in total
Subset: Wrist-only X-rays
Views: Includes PA, AP, and lateral projections
Challenges:
  Class imbalance (more normal cases)
  Variation in image quality and resolution
DATASET

## Slide 6
Label Encoding
converting labels 0 and 1 for negative and positive respectively
Augmentation
Rotation, flipping
Contrast normalization
Gaussian noise
Normalization
Pixel intensities scaled to [0, 1]
Splitting
80% training,
20% validation,
with stratified sampling to maintain class distribution
Modelling
Resnet50 and
EfficientNetB0
Resizing
Standardized to 224×224 for model input
PREPROCESSING PIPELINE
Notes: Preprocessing includes resizing images to 224x224, normalizing pixel values, and applying augmentations like rotations and brightness shifts to prevent overfitting.
We also applied class weights to balance the minority ‘fracture’ class during training.”
🔹 Slide 5: Model Architecture + Hyperparameters (40 sec)
“We used transfer learning with ResNet50 and EfficientNetB0, both pretrained on ImageNet.
We added global average pooling, a dense layer with 64 units, dropout at 0.3, and a sigmoid output layer for binary classification.
We fine-tuned the last 15 layers of each base model.
Common hyperparameters included 25 epochs, a learning rate of 1e-4, and class weights.
We used callbacks like EarlyStopping and ReduceLROnPlateau to stabilize training.”

## Slide 7
DIMENSION DISTRIBUTION IN ORIGINAL DATASET

## Slide 8
MODEL ARCHITECTURE
COMMON HYPERPARAMETERS
Dense units: 64
Dropout rate: 0.3
Learning rate: 1e-4
Epochs: 25
Class weights: {0: 1.0, 1: 1.5}
to address class imbalance (fracture = minority)
Base: ResNet50 pretrained on ImageNet
Top Layers:
    Global Average Pooling
    Dense (64 units, ReLU)
    Dropout (rate = 0.3)
    Output: Dense(1, sigmoid) for binary classification
Fine-tuning: Last 15 layers unfrozen
Optimizer: Adam, Loss: binary_crossentropy, Metric: accuracy
RESNET50 ARCHITECTURE
EarlyStopping: Monitors val_loss, restores best weights if no improvement after 5 epochs
ReduceLROnPlateau: Reduces learning rate by factor 0.2 if no improvement in 2 epochs (min LR = 1e-6)
CALLBACKS
Base: EfficientNetB0 pretrained on ImageNet
Top Layers:
  Global Average Pooling
  Dense (64 units, ReLU)
  Dropout (rate = 0.3)
  Output: Dense(1, sigmoid)
Fine-tuning: Last 15 layers unfrozen
Optimizer: Adam, Loss: binary_crossentropy, Metric: accuracy
EFFICIENTNETB0 ARCHITECTURE

## Slide 9
Metric | ResNet50 | EfficientNetB0
Accuracy | 92% | 88%
Precision (Positive) | 93% | 90%
Recall (Sensitivity) | 87% | 81%
Specificity | 96% | 93%
F1-Score (Positive) | 0.90 | 0.85
AUC-ROC Score | 0.9146 | 0.8707
Interpretation:
  ResNet50 demonstrates stronger overall performance, especially in F1-score and AUC and Specificity.
  EfficientNetB0 offers slightly faster training with comparable precision.
  Both models perform well with literature benchmarks, where AUC typically ranges from 0.87 to 0.93 (Rajpurkar et al., 2017).
PERFORMANCE COMPARISON

## Slide 10
CONFUSION MATRIX
RESNET50 MODEL

## Slide 11
CONFUSION MATRIX
EFFICIENTNETB0
MODEL

## Slide 12
LOSS ERROR GRAGH RESNET50

## Slide 13
LOSS ERROR EFFICIENTNETB0GRAGH

## Slide 14
Discussion
  Both ResNet50 and EfficientNetB0 showed high effectiveness in binary wrist fracture classification using X-ray images from the MURA dataset.
  ResNet50 outperformed across most metrics:
  Accuracy: 92%, Sensitivity: 87%, Specificity: 96%, AUC: 0.9146
  EfficientNetB0 maintained high precision (90%), offering a lightweight and fast alternative suitable for real-time or mobile deployments.
  These results align well with literature (AUC range: 0.87–0.93), supporting their validity.
Limitations
  Grad-CAM failed to produce meaningful fracture localization in several cases due to:
  Subtlety of fractures: Many fractures lack distinct visual boundaries that Grad-CAM can highlight.
  Small lesion size: Grad-CAM’s heatmap resolution is coarse, making it less suitable for fine-grained features like micro-fractures.
  Model focus: The model may have learned spurious correlations (e.g., image artifacts or positioning), diverting attention from the actual fracture site.
  As a result, Grad-CAM interpretability was limited and inconsistent, and not reliable for clinical localization of wrist fractures in our setting
DISCUSSION &LIMITATIONS

## Slide 15
CONCLUSION &
FUTURE WORK
Conclusion
  CNN-based models like ResNet50 and EfficientNetB0 provide high-accuracy, automated wrist fracture detection, with strong diagnostic potential.
  These tools can assist healthcare providers, especially in low-resource or high-demand environments, by improving diagnostic accuracy and reducing radiology workload.
Future work
  Train on datasets with detailed annotations or segmentation masks to guide model attention.
  Incorporate higher-quality, clinically annotated datasets (e.g., bounding box or pixel-level labels).
  Use multi-view X-ray data to improve spatial understanding.
  Explore Vision Transformers (ViTs) or multi-branch CNNs for better spatial feature capture.
  Apply ensemble learning to integrate multiple model decisions for robust performance.
  Validate model performance with radiologist-labeled external data.
  Assess integration potential into PACS systems or radiology workflows in clinical environments

## Slide 16
THANK YOU
Deep learning models like ResNet50 and EfficientNetB0 can serve as effective diagnostic assistants, offering real-time and accurate wrist fracture detection. These models show strong potential for integration into clinical triage systems, especially in environments with limited radiological expertise
QUESTIONS?